| Dokumentennummer / Document Number | 0103-194825 | Fertig gestellt / finalized |
| --- | --- | --- |
| Version / Version | 00 |  |
| Ersteller / Created by | Esber Isshak |  |
| Erstellt Datum / Created Date | 13.11.2019 |  |
| Dokumentenname / Document name | Tender template SOLOASSIST IIS - EN |  |
| Ablageort / File | 1D59A48161FEE43 |  |

| Funktion | Name | Datum | Signatur | Vermerk |
| --- | --- | --- | --- | --- |

| Version / version | Änderung / Changes | Name / name | Datum / Date |
| --- | --- | --- | --- |
| 00 | Erstellung / created |  |  |

Dokumententitel

SOLOASIST IIS ACTIVE ASSISTENCE SYSTEM FOR LAPAROSCOPIC SURGERY

| Pos. | Description | Recommended Quantity |  |
| --- | --- | --- | --- |
| (A) | SOLOASSIST IIS - Field of Application - GENERAL SURGERY - UROLOGY - GYNECOLOGY | - |  |
| (B) | SOLOASSIST IIS ROBOTIC ARM SYSTEM For precise robotic movement and manual positioning of the endoscope during laparoscopic surgery. Mounted direct to the OR Table with an quick connection clamp, suitable for any standard OR Table. SOLOASSIST IIS - Moving Axis: - 3x driven electromechanically - 1x passive universal joint SOLOASSIST IIS - Movement Range: - Zoom: > 400 mm - Rotation: 360° multi-turn - Inclination: +/- 90 ° The system can placed free on the OR table. No fixed position in respect to the Trocar used for endoscope. All control elements are integrated in the SOLOASSIST ARM SYSTEM in surgeons reach: - Button “TP Set” - Button “Manual movement” - LED for TP, Service - Emergency Stop Connection socket for power supply and Voice Control SIIVoice Power supply and cable and power cord (EU) included Rated Voltage: 100-240V ~ / 47-63 Hz Rated Power: 50 VA Weight: 10,4 kg All software included, no external control unit. Transport Trolley included | 1 |  |
| (C) | SOLOASSIST IIS - SYSTEM OVERVIEW 1 Control panel 2 Releasing slider 3 UNIVERSAL JOINT 4 EMERGENCY STOP 5 CLAMP 6 KNURLED NUT 7 STYLUS TIP 8 Handle 9 Quick-fastener for fastening to the operating table 10 Connection VOICE CONTROL 11 Connection POWER SUPPLY | - |  |
| (D) | SOLOASSIST IIS – Voice Control SIIVoice SIIVoice - Voice control system consists of: - Control Unit - Wireless headset - Software included - Connection cable to SOLOASSIST SIIVoice - Key features - Speaker independent - no speech training needed - high responsive - average speech-to-movement time less than 200 ms - precise – more than 95% recognition rate SIIVoice – Functions “Moving the endoscope The Voice Command Set allows precise control over the positioning of the endoscope during surgery. Each command sequence starts with the keyword "SOLO". Once this keyword has been recognized, commands are accepted within 10 seconds. During this period, the "Voice input" indicator lights green. Display Status – Standby Display Status – Active The keyword has been recognized, the VOICE CONTROL is now ready to accept commands, which are then executed by the SOLOASSIST IIS. 10 seconds after the last detected command, the VOICE CONTROL returns to standby mode. SIIVoice - Motion commands of the VOICE CONTROL A motion command consists of direction (LEFT, RIGHT, UP, DOWN) or zoom (IN, OUT), and distance (ONE, TWO, THREE, FIVE, TEN). The direction can be modified to 45° with the command words UPPER and LOWER. The keyword MOVE repeats the last movement. All directions refer to the monitor image. Pay attention to the anatomically correct adjustment of the image horizon. Saving positions Sometimes it may be helpful to memorize a special endoscope position. With the SAVE and RESTORE keywords you can save up to three positions and restore this view when needed. With RETURN it is possible to undo the last movement. SIIVoice - Motion commands “exemple” Desired image change------------------------------- Requiered command Image to the right -------------------------------------------------ONE RIGHT Image to the left -----------------------------------------------------ONE LEFT Image up -----------------------------------------------------------------ONE UP Image down --------------------------------------------------------ONE DOWN Image 45° upwards to the right -----------------------ONE UPPER RIGHT Image 45° upwards to the left ---------------------------ONE UPPER LEFT Image 45° downwards to the right ------------------ONE LOWER RIGHT Image 45° downwards to the left ----------------------ONE LOWER LEFT Overview ---------------------------------------------------------------ONE OUT Detail view ----------------------------------------------------------------ONE IN Repeat movement -------------------------------------------------------MOVE Speed settings With the SPEED keyword followed by ONE, TWO or THREE the speed of the movements of the SOLOASSIST can be adjusted to three pre-set speeds. |  | 1 |
| (E) | SOLOASSIST IIS - STERILE SET consists: - ENDOSCOPE HOLDER SET (1pcs) - STERILE COVER (1pcs) Endoscope Holder SET consists: - Universal Joint - Endoscope Clamp Ø10mm - Endoscope Clamp Ø5mm Sterile Cover – for covering the robotic arm SOLOASSIST IIS - STERILE SET – Key Features The universal joint connects the endoscope clamp with the robotic arm. The Universal Joint is single use SOLOASIST IIS Endoscope Clamp Ø10 mm The CLAMP Ø10 can safely accommodate all endoscopes with an outer diameter of 8 to 11 mm, regardless the view angle of the endoscope (0°, 30°,…) The endoscope clamp is single use SOLOASIST IIS Endoscope Clamp Ø5 mm The CLAMP Ø5 can safely accommodate all endoscopes with an outer diameter of 4 to 6 mm, regardless the view angle of the endoscope (0°, 30°,…) The endoscope clamp is single use |  | 1 SET per procedure 1 per procedure 1 per procedure 1 per procedure 1 per procedure |